## QC Checklist
SCHEDULE M — GMP COMPLIANCE CHECKLIST ¦ QUALITY CONTROL
QC Lab Readiness · Equipment Qualification · OOS/OOT Investigation ¦ Schedule M, Drugs & Cosmetics Rules 1945 (G.S.R.227(E)/2023) ¦ Powered by schedulem.in
▌ MINISTRY OF HEALTH & FAMILY WELFARE, GOVT. OF INDIA ¦ CDSCO / State Licensing Authority ¦ Quality Control Department
Manufacturing Licence No.: |  |  |  | Dosage Form / Product Category: |  |  |  | COMPLIANCE SCORE
Firm Name & Address: |  |  |  | Audit / Inspection Date: |  | 08-Mar-2026 |  | 0 / 58
State Licensing Authority: |  |  |  | QC Head / Authorised Signatory: |  |  |  | ✓ 0
Schedule M Reference: |  | G.S.R. 227(E) dated 28 Mar 2023 |  | Document No.: |  | QA-CHK-QCL-001 Rev.01 |  | ✗ 0
# | Clause Ref. | Checklist Item / Requirement | Guidance Note / Verification Method | Compliance Status | Evidence / Reference Record | Observations & Findings | Corrective Action Required | Responsible / Due Date
6.1 — QC PERSONNEL [Schedule M, Part I, Cl. 6.1 ¦ Drugs & Cosmetics Rules 1945 (Amdt. 2023)]
1 | 6.1.1 | Head of Quality Control is a qualified person (graduate pharmacist / M.Sc. Chemistry or equivalent) independent from production; designation notified to the State Licensing Authority. | Check appointment letter, qualification certificate, Form 28 filing. | Not Checked
2 | 6.1.2 | All QC laboratory analysts hold appropriate qualifications relevant to their assigned tests (chemical, microbiological, instrumental); qualification matrix is documented and current. | Qualification matrix / competency matrix reviewed; records for each analyst on file. | Not Checked
3 | 6.1.3 | Analysts are trained and assessed as competent on all test methods they perform before conducting unsupervised testing; training records with assessment scores are maintained. | Analyst training log; method-specific competency sign-off; assessment score ≥ 80%. | Not Checked
4 | 6.1.4 | GMP and data integrity training (including ALCOA+ principles) is provided to all QC staff at induction and refreshed annually; records are complete. | Training register; data integrity training material; test/quiz records. | Not Checked
5 | 6.1.5 | Annual training plan for QC includes: GMP updates, pharmacopoeial changes, instrument operation, method updates, OOS/OOT handling, and safety. | Annual training calendar approved by Head QC/QA. | Not Checked
6 | 6.1.6 | Personnel performing microbiological testing have received specific training in aseptic technique, media preparation, and environmental monitoring procedures. | Microbiologist-specific competency records; aseptic technique assessment observed. | Not Checked
7 | 6.1.7 | Medical fitness of QC personnel (especially those in microbiology lab) is assessed at recruitment and periodically; exclusion procedure for illness/injury is defined. | Periodic medical examination record; illness exclusion and return-to-work SOP. | Not Checked
8 | 6.1.8 | Deputies are designated for all key QC roles; deputies are trained and their authority is documented. | Deputisation letters; deputy training records. | Not Checked
6.2 — QC LABORATORY PREMISES [Schedule M, Part I, Cl. 6.2 ¦ Separate from production; adequate size & design]
9 | 6.2.1 | QC laboratories are physically separated from production areas; access is controlled; no product manufacturing activity is carried out in the lab. | Floor plan confirms separation; access log or key card control in place. | Not Checked
10 | 6.2.2 | Separate laboratory areas are provided for: physical/chemical testing, microbiological testing, instrument analysis (HPLC, GC, AAS), and reference standard/reagent storage. | Lab layout drawing; designated areas labelled and functionally separated. | Not Checked
11 | 6.2.3 | The microbiology laboratory has classified areas (LAF / BSC cabinet for sterility testing / environmental monitoring); pressure differential, airflow, and temperature are controlled and monitored. | Microbiology lab qualification; environmental monitoring records; airflow visualisation test. | Not Checked
12 | 6.2.4 | Laboratory surfaces (benches, floors, walls) are smooth, impervious, chemical-resistant, and easy to clean; fume cupboards / exhaust ventilation are present where volatile/toxic chemicals are used. | Visual inspection; fume hood certification (face velocity ≥ 0.5 m/s); cleaning SOP. | Not Checked
13 | 6.2.5 | Temperature and humidity in the laboratory are controlled and continuously monitored; records are reviewed and archived; deviations trigger investigation. | Temperature/humidity logger data; alert and action limits defined; deviation investigation records. | Not Checked
14 | 6.2.6 | Adequate and appropriate storage is available for samples (retain, in-test), reference standards (cold chain, desiccator, controlled access), and reagents (flammable, corrosive segregation). | Cold storage qualification; reference standard vault access log; chemical segregation compliance. | Not Checked
15 | 6.2.7 | A dedicated, access-restricted area is provided for storage of narcotic, psychotropic, and other Schedule X / controlled substances used in QC testing. | Narcotics register; double-lock storage; access restricted to authorised persons only. | Not Checked
16 | 6.2.8 | Adequate write-up space, computer terminals, and secure storage for laboratory notebooks / worksheets are available; original data is not recorded on loose sheets. | Bound laboratory notebooks; electronic LIMS/LES with audit trail; data storage SOP. | Not Checked
17 | 6.2.9 | Laboratory waste (chemical, biological, sharps) is segregated, labelled, stored, and disposed of in accordance with applicable regulations; records are maintained. | Waste disposal SOP; waste manifest/records; segregation inspection. | Not Checked
18 | 6.2.10 | Safety equipment (eye wash, emergency shower, fire extinguisher, first aid kit, spill kit) is present, accessible, inspected on schedule, and personnel are trained in its use. | Safety inspection log; training records; eye wash weekly flush record. | Not Checked
6.3 — EQUIPMENT QUALIFICATION [Schedule M, Part I, Cl. 6.3 ¦ IQ / OQ / PQ / Calibration ¦ NABL-traceable]
19 | 6.3.1 | A complete equipment register is maintained for all QC instruments and equipment, including unique asset ID, make, model, serial number, location, qualification status, and calibration due date. | Equipment register in QMS; last reviewed ≤ 6 months; all instruments listed. | Not Checked
20 | 6.3.2 | All critical QC instruments (HPLC, GC, UV-Vis, FTIR, AAS, Karl Fischer, dissolution apparatus, balance, pH meter, particle size analyser) have completed IQ/OQ/PQ; documentation is approved and available. | Qualification dossiers reviewed for 5+ instruments; approved by Head QC and QA. | Not Checked
21 | 6.3.3 | HPLC systems are qualified per USP <1058> / ICH Q2(R1) requirements; SST (system suitability tests) are performed at the start of each analytical sequence and results are within limits. | HPLC qualification report; SST criteria in analytical methods; SST logbook/LIMS records. | Not Checked
22 | 6.3.4 | UV-Vis spectrophotometers are verified for wavelength accuracy, photometric accuracy, stray light, and resolution at defined intervals; results are within pharmacopoeial limits. | UV calibration records using certified filters/standards; frequency ≥ 6-monthly or as per pharmacopoeia. | Not Checked
23 | 6.3.5 | Dissolution apparatus (Apparatus 1 & 2) is qualified per USP <711> / IP requirements including basket/paddle wobble, vessel verticality, rotation speed, temperature, and calibration with prednisone/salicylic acid tablets. | Dissolution apparatus qualification report; mechanical calibration data; calibrator tablet results. | Not Checked
24 | 6.3.6 | Analytical balances are calibrated using certified reference weights (OIML Class E2/F1 or equivalent) and span-checked daily before use; results are logged. | Balance calibration certificate (NABL); daily balance check log; acceptance criteria defined. | Not Checked
25 | 6.3.7 | pH meters are calibrated using at least two certified buffer solutions bracketing the expected sample pH before each use; electrode condition (slope ≥ 95%) is verified and recorded. | pH calibration log; buffer certificate of analysis; slope/asymmetry potential recorded. | Not Checked
26 | 6.3.8 | Stability chambers (ICH zones: 25°C/60%RH, 30°C/65%RH, 40°C/75%RH, refrigerator 2–8°C) are qualified, continuously monitored, and fitted with alarms; temperature mapping is completed. | Chamber qualification report (IQ/OQ/PQ including mapping); continuous data logger records; alarm test records. | Not Checked
27 | 6.3.9 | Autoclaves and hot air ovens used for sterilisation/depyrogenation are validated; cycle validation records (Fo value for autoclave; temperature distribution for oven) are available and current. | Autoclave/oven validation reports; revalidation frequency defined; cycle records for last 3 months. | Not Checked
28 | 6.3.10 | Equipment re-qualification is triggered by: major repair or modification, relocation, performance failures, regulatory requirement, or lapse of qualification period; change control is followed. | Change control SOP; re-qualification trigger criteria defined; recent examples verified. | Not Checked
29 | 6.3.11 | Calibration programme covers all measuring instruments; calibrations are performed by NABL-accredited laboratory or with NABL-traceable standards; calibration certificates are retained. | Calibration schedule; certificates with NABL accreditation number and traceability statement. | Not Checked
30 | 6.3.12 | Out-of-calibration instruments are immediately tagged, removed from service, and an impact assessment is performed on all test results obtained since last valid calibration. | OOC SOP; impact assessment template; review last 2 OOC events for completeness. | Not Checked
31 | 6.3.13 | Computer-controlled instruments and LIMS/software are validated (CSV); access control, audit trails, and backup procedures are implemented and operational. | CSV documentation; audit trail demonstration; backup schedule and restoration test record. | Not Checked
32 | 6.3.14 | Equipment maintenance (PM) is scheduled and completed on time; PM records include activity details, parts replaced, post-maintenance performance check, and sign-off. | PM schedule; PM completion rate ≥ 95% for last 6 months; sample PM records reviewed. | Not Checked
OOS — OOS / OOT INVESTIGATION PROCEDURE [Schedule M ¦ ICH Q10 ¦ 21 CFR 211.192 ¦ WHO TRS 986 Annex 7]
33 | OOS.1 | A written, QA-approved OOS (Out-of-Specification) procedure is in place, covering: Phase I (laboratory investigation), Phase II (full investigation), and reporting requirements. | OOS SOP approved by QA; version current; accessible in lab. | Not Checked
34 | OOS.2 | Phase I laboratory investigation is initiated immediately upon OOS result; analyst reviews: calculations, instrument function, reagent/standard suitability, and sample preparation before invalidating. | OOS Phase I checklist in SOP; Phase I completed within 1 business day; documented. | Not Checked
35 | OOS.3 | Phase I investigation includes: recheck of raw data/chromatograms, verification of instrument SST/calibration status, reagent/standard COA review, and analyst interview; findings documented. | Phase I investigation report template; all steps documented with objective evidence. | Not Checked
36 | OOS.4 | If Phase I does not identify an assignable cause, Phase II investigation is initiated; Phase II involves Production/QA review, reserve sample testing, and additional testing with independent analyst or instrument. | Phase II SOP; cross-functional investigation team; timeline defined (typically ≤ 30 days). | Not Checked
37 | OOS.5 | Hypothesis for each OOS is documented; root cause is identified and classified (laboratory error, process deviation, raw material failure, etc.); conclusion is supported by data. | Root cause analysis template (5-Why / Ishikawa); documented evidence for conclusion. | Not Checked
38 | OOS.6 | Retesting and re-sampling protocols are pre-defined in the OOS SOP; number of retests, acceptance criteria, and statistical treatment are justified and approved by QA. | Retesting justification in SOP; retesting not used to dilute OOS result; QA approval required. | Not Checked
39 | OOS.7 | OOS results are not averaged with passing results unless a laboratory assignable cause is confirmed; passing/failing results are not cherry-picked; all results are reported. | Data integrity check; review OOS register for compliance with averaging rules. | Not Checked
40 | OOS.8 | OOT (Out-of-Trend) procedure is defined; trend analysis is conducted on stability data, process analytical data, and environmental monitoring data; OOT triggers investigation before OOS occurs. | OOT SOP; trend analysis records; statistical method defined (regression, control charts). | Not Checked
41 | OOS.9 | OOS results are reported to QA and production immediately; affected batches are placed under quarantine pending investigation; regulatory reporting requirements are assessed. | OOS notification procedure; quarantine records for affected batches; regulatory reporting SOP. | Not Checked
42 | OOS.10 | CAPA arising from OOS/OOT investigations is tracked to completion; effectiveness is verified; OOS trend data is reviewed at periodic Quality Review meetings. | CAPA log with OOS-linked actions; effectiveness verification records; Quality Review meeting minutes. | Not Checked
43 | OOS.11 | An OOS register / log is maintained with: product, batch, test, OOS value, specification, investigation reference, conclusion, and CAPA status; it is reviewed monthly by Head QC. | OOS register current; last monthly review signed; trend analysis included. | Not Checked
44 | OOS.12 | Laboratory errors identified during OOS investigation are documented, root-caused, and used to update training, SOPs, or analyst competency assessments; repeat errors trigger escalation. | Linked CAPA for lab error; training update evidence; repeat error escalation criteria defined. | Not Checked
LR — LABORATORY READINESS & ANALYTICAL CONTROLS [Schedule M ¦ IP 2022 ¦ ICH Q2(R1) ¦ WHO GMP Guidelines]
45 | LR.1 | Reference standards (primary and working) are sourced from pharmacopoeial authorities (IP/BP/USP) or qualified suppliers; CoA, certificate of assignment, storage conditions, and expiry are documented. | Reference standard register; CoA on file; storage conditions verified; expiry date current. | Not Checked
46 | LR.2 | Working standards are qualified against primary standards; assignment reports with potency/purity determination are available; re-standardisation is triggered on expiry or significant environmental excursion. | Working standard assignment report; potency certificate; re-standardisation SOP. | Not Checked
47 | LR.3 | Reagents, solvents, and culture media are labelled with: name, concentration/grade, preparation date, expiry, and preparer's initials; expired materials are segregated and disposed. | Visual inspection of all reagent/solvent labels in lab; no unlabelled or expired items in use. | Not Checked
48 | LR.4 | Analytical methods used are validated per ICH Q2(R1) (specificity, linearity, accuracy, precision, LOD/LOQ, robustness as applicable); validation reports are approved and available. | Method validation reports for all compendial and in-house methods; reports approved by Head QC. | Not Checked
49 | LR.5 | Compendial methods (IP/BP/USP) are verified for suitability in the laboratory's specific conditions before routine use; verification records are available. | Method verification reports (accuracy, precision, specificity); at minimum for new compendial methods. | Not Checked
50 | LR.6 | Sampling plans and sampling SOPs for raw materials, in-process, finished product, and stability are documented, approved, and followed; sample sizes are statistically justified. | Sampling SOPs; sampling plans referencing IP/ANSI standards; sampling records reviewed. | Not Checked
51 | LR.7 | Retain samples of each batch of finished product are stored under labelled, defined conditions for the required retention period (at least 1 year beyond expiry or as per local regulations). | Retain sample register; storage conditions verified; retention period compliant with Schedule M. | Not Checked
52 | LR.8 | Stability programme covers: accelerated (40°C/75%RH), intermediate (30°C/65%RH), and long-term (25°C/60%RH or as per label) conditions; on-going stability programme is active for each product. | Stability protocol for each product; stability schedule; current stability data reviewed. | Not Checked
53 | LR.9 | Laboratory notebooks / worksheets are completed in real time, in ink, with no erasures (errors struck through and initialled); all raw data (printouts, chromatograms) is attached and identifiable. | Data integrity audit of 5 recent worksheets; ALCOA+ compliance verified. | Not Checked
54 | LR.10 | Electronic data from instruments (HPLC, GC, UV) is stored securely with audit trails enabled; no unapproved deletion or alteration of raw data files; back-up is scheduled and tested. | LIMS/CDS audit trail demonstration; backup log; data access control matrix. | Not Checked
55 | LR.11 | Environmental monitoring programme for cleanrooms/microbiology lab includes: viable air sampling, non-viable particle counting, surface sampling, and personnel monitoring; alert and action limits are defined. | EM SOP; EM schedule; last 3 months EM data reviewed; OOL investigation records. | Not Checked
56 | LR.12 | Water for pharmaceutical use (Purified Water / WFI) is tested for chemical and microbiological quality at defined sampling points and frequencies per the approved water monitoring programme. | Water SOP; sampling schedule; last 3 months water test results; trend analysis. | Not Checked
57 | LR.13 | A market complaint testing procedure is in place; QC lab can test complaint samples against approved specifications; retain samples are used for comparison. | Complaint testing SOP; recent complaint test records; link to CAPA. | Not Checked
58 | LR.14 | Annual Product Quality Review (APQR) data contribution from QC (test result trends, OOS count, stability data, reference standard usage) is prepared and provided to QA on schedule. | APQR contribution records; last APQR chapter from QC reviewed. | Not Checked
Head QC Signature: _____________________________ Date: __________ |  |  |  |  | Head QA Signature: _____________________________ Date: __________
QA-CHK-QCL-001 ¦ Rev: 01 ¦ 08-Mar-2026 ¦ Schedule M, Drugs & Cosmetics Rules 1945, G.S.R.227(E)/2023 ¦ © schedulem.in ¦ CONFIDENTIAL

## Audit Summary
SCHEDULE M QC AUDIT — SUMMARY DASHBOARD
Quality Control · Lab Readiness · OOS Investigation · Equipment Qualification ¦ © schedulem.in
▌ Auto-calculated from QC Checklist tab ¦ Schedule M (G.S.R.227(E)/2023)
| TOTAL ITEMS | COMPLIANT | NON-COMPLIANT | PARTIAL | NOT CHECKED
| 58 | 0 | 0 | 0 | 58
| items | items | items | items | items
OVERALL SCHEDULE M QC COMPLIANCE RATE
0%
| Section | Total | Compliant | Non-Compliant | Partial | Score %
| OOS / OOT Investigation | 12 | 0 | 0 | 0 | 0%
| Lab Readiness & Controls | 14 | 0 | 0 | 0 | 0%
STATUS LEGEND
| ✓ Compliant | Fully meets Schedule M / pharmacopoeial requirement; evidence on file.
| ✗ Non-Compliant | Requirement not met; corrective action mandatory before next inspection.
| ~ Partial | Partially meets requirement; improvement required within agreed timeframe.
| N/A | Not applicable to this dosage form, product, or lab scope.
| Not Checked | Not yet reviewed during this audit cycle.
QA-CHK-QCL-001 Rev 01 ¦ 08-Mar-2026 ¦ Schedule M, Drugs & Cosmetics Rules 1945 ¦ © schedulem.in ¦ CONFIDENTIAL